



**HED Internal Application Error, Patient is on a ventilator-** Nurse is unable to save charting 11/1/13/kw


Cause: The ventilator is sending DAS a text value to a data entry field where HED expects a numerical value.


If this error is seen in HED and the patient is on a ventilator:  
  


Ask the nurse to check the HED flowsheet for a value similar to these: STANDB\* and ADULT.  




**Workaround is to clear that value from the data entry field or uncheck the box if checked.**

Cleared STANDB\*

Unchecked STANDB\*  


**Fix is to have the nurse remove the VueLink module from the monitor, wait 10 seconds and reconnect.   
If that does not correct the problem, Respirator y Therapy will need to restart the ventilator.**