Monitored Vital Signs- 10 Tips for Making DAS and DAS for Vents Work for You (and the Users)

1) Users do not call the system "DAS"; they usually call it "Monitored Vital Signs" when they call into the Help Desk. DAS is an abbreviation for “Data Acquisition System”

2) Floors that use DAS: SICU (3N/C); CVICU (5N); Neuro ICU (5S); MICU (7S); Surgical Stepdown (9N); Trauma (10N); 10S (on s/p kidney transplant pts only); Burn (11S); Myelo-if monitoring is ordered (11N); NICU (4NI1-2, A-J); PCCU (5PCU)

3) To select all the measurements for a particular time point, select the checkbox next to time. To open up all of the other available documentation fields, click the black triangle next to the time. Everything that is imported can be changed by the RN. Just click on the number to change and type over it

4) If DAS is not working on ANY patient on the entire unit, the server is down. Ask the Help Desk to refer ticket to the Biomed call person as high priority and notify HED Application Support On-Call Person.

5) If DAS is not working at all on only one patient, make sure the patient admitted to the bed matches the patient admitted in the monitor

6) When a patient is moved from the bed for any reason in Medipac (leaves the floor for radiology, surgery, moves to another bed on the same unit, etc), all DAS data is lost. Tell RNs to save their most recent VS right before they move their patient

7) When only certain measurements are not working in DAS that means the parameter is not set correctly on the Philips monitor. Use the chart below to determine how the monitor should be set for each measurement.

8) For hemodynamic measurements, certain VueLink modules (the cartridges that plug into the bottom of the Philips monitors) are needed for measurements to be available in DAS. In all of the ICUs another freestanding monitor is used to measure hemodynamics (Vigilance monitors). Vigilance monitors are then slaved into the Philips monitors and hemodynamic parameters become available in DAS. Once the Vigilance monitor is slaved into the Philips monitor, the RN should not have to adjust any settings on the Philips monitor for hemodynamic measurements to work in DAS. Not all hemodynamic measurements are available on every unit. Some units have older VueLink modules that do not support CO, CI, and Core Temp measurement.

9) For Vents, a certain VueLink module is required. Currently, only vent settings for Servo I vents are available. The modules are attached to each Servo I vent and the Respiratory Therapy Department is responsible for the modules. Once the module is plugged into the Philips monitor, the RN and/or RT should not have to adjust any settings on the Philips monitor

10) Some Vent measurements do not come over from Philips with the same units of measure that the RNs traditionally document them in (e.g., Inspired Minute Volume is usually measure in ml; however with DAS for Vents it is now documented in L). This can cause some frustration for some RNs.

| Charting Parameter | Monitor Setting |
| --- | --- |
| All VS Except NICU |
| Heart Rate/DAS | HR |
| ABP and ABP mean | ABP |
| Art BP and ART BP mean | ART |
| NIBP and NIBP mean | NIBP |
| Respirations | Resp |
| O2 Sat | SpO2 |
| Hemodynamics (Vigilance monitors slaving into Philips) |
| Core Temp in C | CVICU only |
| PAP and PAP mean | All ICUs |
| CVP | All ICUs |
| CO | CVICU only |
| CI | CVICU only |
| Neuro Parameters |
| ICP | ICP |
| ICP#2 | UVP |
| Other |
| Bladder Pressure | Pressure #2 |
| Vents (Module controls Philips display) |
| Total Resp Rate | --- |
| Inspired Tidal Volume | --- |
| Inspired Minute Volume | --- |
| Tidal Volume Expiratory | --- |
| Peak Pressure | --- |
| Mean Airway Pressure | --- |
| PEEP Measured | --- |
| Inspired Oxygen Concentration | --- |
| NICU VS |
| Heart Rate | HR |
| Respirations | Resp |
| ABP and ABP mean | ART (used for both UACs and PALs) |
| NIBP and NIBP mean | NIBP |
| O2 Sat #1 | SpO2 |
| ICP | ICP |